Learning Manual for Assessment of

the Virtual Anesthesia Machine

Version VR 1.1

7/13/06

Learning Manual for Assessment of

the Virtual Anesthesia Machine

Version VR 1.1

7/11/06

During today’s session, you will be learning about how an anesthesia machine works by using a computer simulation of a typical anesthesia machine. The training involves four stages:

- Read an overview of the purpose, function, and structure of the anesthesia machine, designed to give you a framework for understanding the material that follows

- Take a tour of the main subsystems of the machine, and the components and controls of each system, through the simulation

- Focusing on some of the subsystems, learn the answers to several specific questions about the operation of the machine and control of anesthesia, with step-by-step instructions for using the simulation to answer them

- Try to answer several other questions by using the simulation

Tomorrow, we will try to assess what you learned about the machine, and how well you learned it.

1. Overview of anesthesia and the anesthesia machine

General Anesthesia

General anesthesia is the induction of a balanced state of unconsciousness, accompanied by the absence of pain sensation, and the paralysis of skeletal muscle over the entire body. It is induced through the administration of anesthetic drugs and is used during major surgery and other invasive surgical procedures.

General anesthetics may be gases or volatile liquids that evaporate and are inhaled along with oxygen and other atmospheric gases. The amount of anesthesia produced by inhaling a general anesthetic can be adjusted rapidly, if necessary, by adjusting the anesthetic-to-oxygen ratio that is inhaled by the patient.

Functions of the Anesthesia Machine

- Get gases from supply lines

- Measure their pressure and flow, and load them with anesthetic vapors

- Present them to the patient for breathing

- Breathe for the patient, if necessary

- Maintain an appropriate mixture of gases being breathed, adding oxygen and removing carbon dioxide as needed

- Provide an exit route for gases

Main subsystems of the Anesthesia Machine:

- The gases are supplied by a high pressure system.

- An adjustable low pressure system reduces the pressure, and mixes the gases along with the anesthetic vapors.

- Since paralysis is a consequence of general anesthesia, breathing is usually controlled by a manual breathing system, in which the operator squeezes a bag to deliver gases to the patient, or…

- by a mechanical breathing system, in which a bellows is alternately filled with gases, and emptied into the airway.

- The breathing circuit connects the patient to the anesthesia machine, of course, and controls the flow of gases during inhalation and exhalation.

- Finally, the exit route for gases is provided by the scavenging system.

The Virtual Anesthesia Machine: A “virtual reality simulation”

In learning about the function of the anesthesia machine, it is important to be able to visualize what the various components and controls of the system look like, where they are located, and learn how the different subsystems operate to control the flow of gases. The Virtual Anesthesia Machine was intended to give you that “virtual reality” view of the operation of the machine. The ability to see the machine as you learn about it helps you build a “mental model” of the machine, and understand how it does what it does. It also helps you understand what can go wrong with the machine, and respond to unexpected faults.

Because of the importance of visualization,, be sure that you take advantage of this aspect of the simulation, carefully considering the consequences of your actions on machine function, both locally within that system, and globally in other systems.

In the exercises that follow, you will learn more about some of the components and functions of the anesthesia machine by using the Virtual Anesthesia Machine (VAM) simulation to answer a set of questions about three of the subsystems that are central to its function:

- The breathing circuit

- The manual ventilation sub-system

- The mechanical ventilation sub-system

2. Introduction to the VAM and its components

We begin with a “grand tour” of the anesthesia simulation to familiarize you with the location of the major subsystems, and the controls and structures of those subsystems. As you work through the tour, try to imagine how each subsystem and component might operate in conjunction with the other parts of the machine.

The major subsystems:

____ Click on Help Using VAM (in the white box in the lower left of the screen):

____ Click on the blue arrow to see each of the component subsystems:

- High pressure system. For both Oxygen (O2, GREEN) and Nitrous Oxide (N2O, BLUE), note (left side) there are two sources – pipelines from the wall, and cylinders as back-ups. On the right, note the high pressure gauges, and a small green button for flushing high-pressure oxygen into the system.

- Low Pressure system. Includes controls for flow of these gases, and the anesthetic agent (in the two vaporizers).

- Circle System Breathing Circuit. Includes paths for inhalation (upper left valve), exhalation (lower right valve), and the CO2 absorber for “scrubbing” or absorbing carbon dioxide from the gas. Why is this called a circle system?

- Manual ventilation system. Basically a squeeze bag, and a valve that keeps the gas pressure from getting too high during manual ventilation (the Adjustable Pressure Limiting, or APL, valve). Note the position of the ventilation selector switch (on the top of the CO2 absorber). This controls whether ventilation is under manual or mechanical (automatic) control.

- Mechanical ventilation system. Note the large bellows in the middle. As this is compressed, gas is pushed into the lungs; then, as the lungs deflate, it’s refilled with gas from the lungs and from some fresh gases from the low pressure systems

- Scavenging system. There’s a constant, low-pressure suction to remove excess gases from the system. The waste gas scavenging bag fills up during exhalation, and deflates during inhalation.

Interactive User Controls

____ Click the green arrow to see each control. The controls that you will be using in this lesson are highlighted in red. Consider the function of each in the machine and try to imagine its operation:

- Oxygen cylinder supply. Contains enough oxygen for a limited amount of respiration. (Time to exhaustion of the cylinder depends on the consumption rate, which can be very variable. A full cylinder contains 660 L of O2).

- Nitrous Oxide cylinder supply.

- Oxygen pipeline supply. This is delivered from a remote central supply (actually liquid O2 at UF).

- Nitrous Oxide pipeline supply.

- Nitrous Oxide flowmeter knob. In the simulation, you control gas flow by a left-click-and-hold, then drag the knob counterclockwise to increase flow. Note the position of the “bobbin” (small red float) in the flow tube. The higher the bobbin, the greater the flow.

- Oxygen flowmeter knob. Note the position of the bobbin for Oxygen. Is it the same as for the Nitrous Oxide?

- Vaporizer dial. The vaporizer contains liquid volatile anesthetic, the dial controls how much is vaporized and added to the gas mixture.

- Oxygen flush valve. This valve allows oxygen to flow directly to the patient, bypassing the vaporizer and flowmeters. It is used to quickly increase the amount of oxygen and decrease the anesthetic concentration in the system.

- Selector switch. Selects manual or mechanical ventilation. In the illustration, the switch is set for mechanical ventilation.

- APL (Adjustable Pressure Limiting) valve. Can be fully open, fully closed, or set to open at a threshold pressure; effective during manual ventilation only.

- Manual breathing bag. What will happen when the bag is squeezed if the APL is closed? What if it’s wide open? What do you think would happen when the bag is squeezed but the Selector Switch is set for mechanical versus manual ventilation?

- Ventilator switch. Turns on mechanical ventilation, compressing the bellows at a selected breathing rate.

- Wall vacuum plug. Provides a vacuum source and removes the scavenged gases to prevent room pollution.

- Scavenging adjustment valve. Controls the outflow created by the Scavenging vacuum.

C. Other Key Components

- Airway pressure gauge. Given its placement between the inspiratory valve and the lungs, what would happen to the pressure if the patient tried to inhale, or exhale, with the inspiration and expiration valves closed?

- Inspiratory valve. Opens up when the pressure in the airway is less than the pressure on the other side of the valve (as during inhalation, or inspiration); The valve is unidirectional, and gases can’t flow in the opposite direction.

- Expiratory valve. Also unidirectional, but in the opposite direction. Opens when the pressure in the airway is greater than the pressure on the other side of the valve. Try to visualize their operation during respiration.

- CO2 absorber.  “Scrubs” the gas of carbon dioxide during inspiration.

- Pressure gauges for each of the four possible sources of gas

D. Interactive Ventilator Settings.

These control various aspects of the mechanical ventilation sub-system:

- Inspiratory Pressure Limit. Prevents over-inflating of the lungs during mechanical ventilation. Current standards set this at a default of 40 cm H2O.

- Inspiratory pause. During mechanical ventilation, a user-adjustable pause may be added between active inflation and the start of exhalation

- Frequency of breathing, in breaths per minute.

- Tidal volume, in mL, is the total volume of gases inhaled during inspiration.

- The I/E ratio is the proportion of time spent inhaling (Inspiration, I) and exhaling (Expiration, E). Think of your spontaneous breathing cycle. What seems to be the “natural” ratio? Note that there is a short time between exhalation and the next inhalation; this is called the end-exhalation period, and is included in the E interval, as you’ll see in the simulation.

- Common gas outlet checkvalve. Prevents retrograde flow from the circuit (especially if the circuit is at high pressure) back into the vaporizer which could lead to higher than intended volatile anesthetic concentrations and overdose.  (Not all anesthesia machines have this valve.)

E. Other major components (not shown)

- Oxygen failsafe. Automatically shuts off flow of N2O if the O2 supply pressure drops below a threshold, so that a hypoxic gas mixture is not delivered.

- Pressure regulators. One for each of the two input gases, steps the pressure down from high to low pressure systems.

- Low oxygen supply pressure alarm. Warns that the oxygen supply pressure has fallen below a set threshold and will soon fail or has failed.

- Ventilator proportional flow control valve.  Part of the mechanical ventilation system. Opens during mechanical inspiration to pressurize and drive the bellows down.

- Ventilator Pressure Relief (VPR) valve.  Part of the bellows system. Also called the “spill” or the “pop-off” valve). Allows those excess gases to be discharged, once exhalation is at an end, and the bellows have re-filled.

- Scavenging positive and negative pressure relief valves. As the names imply, keeps the pressure within the scavenging system within a designated range.

____ Click the X in the upper right corner of the window to close the “Help using VAM” module.

3. Exploring the Virtual Anesthesia Machine

In each of the sections below, we begin with a question about the function of the machine, and using the simulation, demonstrate and explain the answer, addressing other questions that can be illuminated by the VAM along the way.

As you read each of the questions in italics, try to answer it by first looking at the simulation, and considering the “dynamics” of gas flow and the operation of relative components, before reading the answer.

Be sure you understand the explanation before moving on to the next question. The controls you operate on are indicated in screen shots.

A Preview of the Exploratory Questions.

Here are the five main questions that you’ll be answering during the tutorial:

Question 1: elimination of CO2. Are the gases exhaled by a patient “scrubbed” of CO2 before entering the bellows during mechanical ventilation?

Question 2: changes in fresh gas flow concentration. You are anesthetizing the patient with a concentration of 30% oxygen (O2) mixed with 70% nitrous oxide (N2O). In preparation for emergence from anesthesia back into consciousness, you turn off the N2O flow to set the O2 concentration to 100%. Is your patient now breathing 100% O2?

Question 3: elimination of excess gases.  What is the exit for gas during the inspiratory phase of mechanical ventilation?

Question 4: flushing with oxygen.   Why should you not flush during mechanical inspiration?

Question 5: The APL valve.  The adjustable pressure-limiting (APL) valve is closed after switching to mechanical ventilation for maintenance of anesthesia. Switching back to manual ventilation for emergence from anesthesia, you forget to adjust the APL valve, and squeeze the manual ventilation bag. What happens?

A roadmap and some reminders.

Before beginning, there are some basic aspects of the simulation that are worth pointing out:

- The system can always be reset to the start conditions by clicking the word RESET at the bottom left  (See the screen shot below)

- You can control the components and functions described below by pointing to the component and left-clicking (and sometimes dragging) the icon for the component.

[

Question 1: elimination of CO2. Are the gases exhaled by a patient “scrubbed” of CO2 before entering the bellows during mechanical ventilation?

Demonstration using VAM Simulation:

____  Click “Reset” to start simulation afresh

____  Point to the O2 flowmeter control knob to enlarge it, then click-and-hold, and drag it counterclockwise until the O2 bobbin inside is about halfway up the tube.

What does this do?

Opening the valve increases the flow of O2 from the supply line into the breathing circuit.

Where does it wind up?

It depends. For example, If mechanical ventilation is selected, but not on, the O2  flows “backward” through the  CO2 absorber, past the bellows and into the scavenger system

____  Click on the ventilator on/off switch icon to turn on the mechanical ventilation system. The patient is now being automatically ventilated by the mechanical system.

Notice what’s going on during inspiration, as opposed to exhalation, as you scan through the simulation and consider the flow of gases. Study it until you feel you have some sense of what’s happening. Then verify the following, trying to answer each question by looking at the VAM before reading the answer:

During inspiration:

What powers the bellows?

O2 flows into the bellows casing from the high pressure pipeline. The bellows are compressed, forcing gases in the breathing circuit to flow through the plumbing towards the lungs. Fresh O2 from the low-pressure line can now flow into the lungs as well.

Do the gases from the bellows flow directly into the lungs?

No, they are pushed through the CO2 absorber first.

Why can’t the gases from the bellows flow directly to the lungs?

The exhalation valve is unidirectional, and it’s closed during inhalation.

What happens in the CO2 absorber?

When  molecules of CO2  enter the absorber, they’re absorbed by chemical granules; the other gases pass through.

Where does the gas go from there?

The increased pressure opens the inspiratory valve, gas flows into the lungs, and they expand.

Is most of the inhaled oxygen coming from the bellows and absorber, or from the O2 low pressure pipeline?

Even with the flowmeter halfway open, most of the gases are being rebreathed. That’s why this is called a rebreathing ventilation system.

During exhalation:

What powers the expansion of the bellows?

The lungs deflate due to the elastic recoil of the chest wall, and the exhaled gases flow back into the bellows.

What path do the exhaled gases follow?

Gases flow directly through the exhalation valve back into the bellows without passing through the CO2 absorber.

What prevents exhaled gases from backflowing into the CO2 absorber during exhalation?

Note that during exhalation, the inspiration valve is closed, diverting the fresh O2 flow from the flowmeter to flow retrograde (“backwards” through the CO2 absorber and towards the bellows).  So the exhaled gases take the path of least resistance, to the bellows.

What is the flow path of the fresh gas from the O2 flowmeter during expiration?

Since the inspiratory valve is closed, but O2 is still flowing, it’s forced “backwards” through the CO2 absorber, as we just learned, where it joins with the exhaled gases, to the bellows and, if O2 flow is set very high, out to the scavenger system

What happens to the O2 that compressed the bellows during inhalation?

The ventilator exhalation valve opens, venting the drive gas in the bellows casing to the room, and allowing the bellows to refill.

Review of Question 1: elimination of CO2.  Are the gases exhaled by a patient “scrubbed” of CO2 before entering the bellows during mechanical ventilation?

Answer: No.

Explanation: During mechanical ventilation, gases exhaled by the patient flow directly into the bellows. The CO2 absorber is located so that these exhaled gases are only scrubbed clean of CO2 when they are being redirected to the patient, during inspiration. (During exhalation, there’s only a slow “retrograde” flow of fresh O2 through the CO2 absorber.) The rationale for this is that scrubbing exhaled gases that may then be dumped into the scavenging system would prematurely exhaust the CO2 absorbent, and be wasteful.

Question 2: changes in fresh gas flow concentration. You are anesthetizing the patient with a concentration of 30% oxygen (O2) mixed with 70% nitrous oxide (N2O). In preparation for emergence from anesthesia back into consciousness, you turn off the N2O flow to set the O2 concentration to 100%. Is your patient now breathing 100% O2?

Demonstration using VAM:

____  Click “Reset” to start simulation afresh

___  Click and drag the N2O flowmeter knob icon counterclockwise until the top of the N2O bobbin is near the top of the tube.

What happens to the flow of N2O from the pipeline?

It’s immediately mixed with the fresh O2 gas described earlier, so with mechanical ventilation selected and the ventilator switch off, it flows slowly back through the CO2 absorber and into the scavenging system.

Is there a change in fresh O2 flow? Why?

You may have noted that as you increased the NO2 flow past a certain point, the O2 flowmeter opened up along with it. The two are linked to prevent having the O2 concentration fall below about 20%.

____  Click and drag the O2 flowmeter knob icon counterclockwise until the top of the O2 bobbin is between the first and second mark on the left of the tube.

____  Click on the ventilator on/off switch icon to turn on the mechanical ventilation system.

____  Wait about 30 seconds to let the fresh gas mixture reach the lungs and bellows. Note the Inspired Oxygen Concentration monitor at the top center part of the simulation screen. It should stabilize at about 30% O2.

____  Click and drag the N2O flowmeter knob icon fully clockwise so that there is no more N2O flowing in the N2O flowmeter.

____  Click and drag the O2 flowmeter knob icon fully counterclockwise until the bobbin is near the middle of the flowmeter tube.

Does the residual N2O get expelled from the system, or does it get rebreathed?

N2O molecules remain in the circuit for a fairly long time after the N2O flowmeter has been reduced to zero. Note how slowly the inspired oxygen % concentration changes.

How long do you think it would take for the N2O to clear out completely?

This is a function of the “time constant” of the system at these settings.  (The time constant would be the time it would take for the system to reach 63% of the step change - for example, the time to reach a concentration of about 31% N2O if you were going from 0% to 50% N2O. To clear out completely would require 3 time constants in general.)

Will this time constant be greater if the flow of fresh gas was greater?

No, the time constant will be shorter because the N2O will be cleared faster by the higher flow of fresh gas.

Review of Question 2: changes in fresh gas flow concentration. You are anesthetizing the patient with a minimal flow of nitrous oxide (N2O) mixed in with the fresh O2. In preparation for emergence from anesthesia, you turn off the N2O flow. Is your patient now inhaling gas with 0% N2O?

Answer: No, not for while.

Explanation: Breathing circuits have a “time constant” whereby a change in the set gas composition does not immediately result in a corresponding change in the inspired gas composition. This time constant is influenced by, among other things, the magnitude of the fresh gas flow and the volume of gases in the breathing circuit.

Question 3: elimination of excess gases.  What is the exit for gas during the inspiratory phase of mechanical ventilation?

____  Click “Reset” to start simulation afresh

____  Click and drag the O2 flowmeter knob icon until the bobbin is halfway up the tube.

____  Click on the ventilator on/off switch icon to turn on the mechanical ventilation system.

What is the exit route for gas during mechanical ventilation?

If the selector switch is set to mechanical ventilation, there’s only one place where excess gas can reach the scavenging system below. It’s through a valve at the bottom of the bellows. This valve is called the ventilator pressure relief valve, or VPR valve (or sometimes the “spill” valve).

How does the VPR valve behave during inhalation?

At the end of inspiration, the outer chamber of the bellows is filled with “drive gas” from the high pressure O2 supply. The VPR valve is pressurized shut by the drive gas during inspiration, and stays shut until the drive gas is released to the room.

When does the VPR valve open during exhalation?

The VPR valve opens at or near the end of exhalation, when the bellows has completely refilled.

When does the VPR valve close during inspiration?

The VPR valve closes at the moment inspiration begins.

Why does the VPR valve need to be closed during inspiration?

If it were open during inspiration, the bellows gases could just escape into the scavenging system.

Review of Question 3: elimination of excess gases. What is the exit for gas during the inspiratory phase of mechanical ventilation?

Answer: There is none.

Explanation: The ventilator pressure relief valve is the outlet for gases during mechanical ventilation, but it only relieves gas during exhalation. It is pressurized shut by the drive gas during mechanical inspiration.

Question 4: flushing with oxygen. Why should you not flush during mechanical inspiration?

Demonstration using VAM:

____  Click “Reset” to start simulation afresh

____  Click and drag the O2 flowmeter knob icon until the bobbin is halfway up the tube.

____  Click on the ventilator on/off switch icon to turn on the mechanical ventilation system.

What routes can the high pressure O2 gas follow when the O2 flush valve is pressed? How will this depend on whether the VPR is opened or closed?

Pressing the O2 flush valve allows high-pressure O2 into the breathing circuit, so it will follow the same routes that fresh gas does. During expiration, the VPR valve opens, and the high pressure gas can escape to the scavenging system.

___  Click on the airway pressure gauge icon to enlarge the pressure gauge, so you can read the pressure on the gauge. The inner part of the gauge displays pressure in terms of cm of H2O.

What is the maximum airway pressure during ventilation, and when is it reached?

Watch how the pressure increases steadily during inspiration, reaching a peak of about 18 cm H2O.

____  Press the oxygen flush valve during exhalation.

What route does the O2 take from the high pressure line when the O2 flush valve is pressed during exhalation?

The inspiratory valve in the airway is closed, so it can’t go there. The route is “backwards” through the CO2 absorber, a bit into the expanding bellows, and then, when the VPR valve opens (see question 3), on to the scavenger system.

What is the maximum airway pressure with a flush during exhalation?

Note that as you press the flush valve during exhalation, it has no effect on the (dropping) airway pressure. Why is that?

The  expiration valve is unidirectional, so the flushed O2 gas can’t flow into the airway, and instead takes the route outlined above.

____  Press the oxygen flush valve during inspiration

What route does the O2 take from the high pressure line when the O2 flush valve is pressed during inspiration?

Since the VPR is pressurized closed, the O2 flows into the lungs along with the gases coming from the bellows.

What is the maximum airway pressure with a flush during inspiration?

As you press the flush valve during inspiration, the pressure increases rapidly (this is a high pressure line, after all). It peaks at about 40 cm of H2O, but then quickly dissipates (does not dissipate if you keep pushing the O2 flush).

Why does it peak at 40 cm H2O, rather than continue to increase and create hyperinflation?

Look at the setting of the Inspiratory Pressure Limit (IPL) setting on the upper right.

Review of Question 4: flushing with oxygen. Why should you not flush during mechanical ventilation?

Answer: Because the O2 flush gas will flow to the lungs and hyperinflate the lungs until the inspiratory pressure limit (if properly set) is reached, risking barotrauma.

Explanation: An O2 flush valve delivers about 1 L/sec into the breathing circuit – a lot of gas. With the VPR valve shut during mechanical inspiration, the gas can only wind up in the lungs, raising the pressure to potentially dangerous levels. The default IPL setting on some actual machines may be set anywhere from 40 cm to 60 cm H2O or greater. So, flush when the bellows are moving up!

Question 5: The APL valve. The adjustable pressure-limiting (APL) valve is closed after switching to mechanical ventilation for maintenance of anesthesia. Switching back to manual ventilation for emergence from anesthesia, you forget to adjust the APL valve, and squeeze the manual ventilation bag. What happens?

Demonstration using VAM:

____  Click “Reset” to start simulation afresh

____   Click on the APL valve icon three times until it is fully closed (pressed against the console, with only the cap visible).

Can an open APL valve cause a “leak” during mechanical ventilation?

Note the location of the APL valve within the manual ventilation circuit. It’s located below the switch that selects for manual or mechanical breathing systems, and so is part of a second “exit route” for gases that only functions during manual ventilation. Its setting during mechanical ventilation is irrelevant, at least in this version of the anesthesia machine. (In others, it’s placed above the selector switch and creates a leak if left open during mechanical ventilation).

____  Click and drag the O2 flowmeter knob icon counterclockwise until the O2 bobbin is fully up at the top of the tube:

____  Click on the selector knob icon to select manual ventilation (four-o-clock position. Look at the position of the inspiratory and expiratory valves in the breathing circuit.

Does the airway pressure begin to increase even before you’ve squeezed the bag?

Watch what happens as the bag starts to inflate from the gases being delivered by the O2 flowmeter input. At first, the airway pressure stays low; but then as the bag starts to swell and pressure builds, the inspiratory valve opens, allowing gas into the lungs.

____  Click on the airway pressure gauge icon to enlarge the gauge.

How high can the airway pressure get? Watch as the gauge climbs to about 40 cm H2O and plateaus there.

____  Click on the manual ventilation bag to squeeze it.

What happens to airway pressure?

Momentarily, it climbs to about 70 cm H2O, but quickly returns to about 40 cm H2O.

Is this because of the Inspiratory Pressure Limit setting on the ventilator (remember the setting at the upper right)?

No, the IPL setting only affects the mechanical ventilation system. With manual ventilation, it’s the compliance of the squeeze bag, which, at 40 cm H2O, can increase a lot in volume with no change in pressure.

What happens to the lungs?

40-50 cm H2O is enough to cause some barotrauma, stylized here with reddened tissue.

____  Click on the APL valve once to partially open it.

What happens to airway pressure now?

Pressure is released, and excess gas exits to the scavenger system. Pressure drops to about 16 cm H2O, which is the pressure setting

____	Give the manual ventilation bag a couple of squeezes.

What happens to airway pressure now?

From the resting, APL-controlled pressure of about 16 cm H2O, it peaks at about 24 cm H2O during inspiration, then drops back down during exhalation. You can see the inspiratory valve, then the expiratory valve, open and close in turn during manual ventilation. This is how the system functions normally.

Review of Question 5: The APL valve. The adjustable pressure-limiting (APL) valve is closed after switching to mechanical ventilation for maintenance of anesthesia. Switching back to manual ventilation for emergence from anesthesia, you forget to adjust the APL valve, and squeeze the manual ventilation bag. What happens?

Answer: It gradually increases until it plateaus at about 40 – 50 cm H2O.  But you can squeeze to generate higher pressures up to 70 cm H2O.  A totally closed APL valve won’t relieve until pressure reaches 70 cm H2O.

Explanation: The function of the APL valve is to limit airway pressure during manual ventilation. In the simulation, it’s set for about 16 cm H2O. Fully closed, it will only let gas flow through it when pressure reaches about 70 cm H2O. The breathing bag can continue expanding with minimal change in pressure once we reach 40 – 50 cm H2O.

This completes the training tutorial on the anesthesia machine.

While you’re not yet ready for an internship in anesthesiology, you’ve learned some of the basics of operation of the standard machine used worldwide during general anesthesia, and some of the things that need to be monitored during its use.

This is a good time to review again the questions we’ve answered during the tutorial. Just read through the list below, and see if you can answer the questions by remembering those conditions in the simulation.

Question 1: elimination of CO2. Are the gases exhaled by a patient “scrubbed” of CO2 before entering the bellows during mechanical ventilation?

Question 2: changes in fresh gas flow concentration. You are anesthetizing the patient with a concentration of 30% oxygen (O2) mixed with 70% nitrous oxide (N2O). In preparation for emergence from anesthesia back into consciousness, you turn off the N2O flow to set the O2 concentration to 100%. Is your patient now breathing 100% O2?

Question 3: elimination of excess gases.  What is the exit for gas during the inspiratory phase of mechanical ventilation?

Question 4: flushing with oxygen.   Why should you not flush during mechanical inspiration?

Question 5: The APL valve.  The adjustable pressure-limiting (APL) valve is closed after switching to mechanical ventilation for maintenance of anesthesia. Switching back to manual ventilation for emergence from anesthesia, you forget to adjust the APL valve, and squeeze the manual ventilation bag. What happens?